## Detector Prototype data avg.
Data: 19 Oct 1992 |  | Thermal drift, Linearity and Vout vs Load of zero bias Schottky Diode: Used in 4503, 4506, 70122 & 70130
Redrawn: 18 Apr 2006
|  |  | Diode ID: Avg of Wht 1-Wht 5 prototypes |  |  |  | Rvideo=Rload*((V10Mohm/Vload)-1) |  | Rvideo= Rdiode + 300
Power Level | Load resistor | Temperature (Degrees C): 23 deg data is the Output Voltage measured. At other Temperatures the data given is the drift relative to 23 deg in dB using the diode Linearity- not a simple "10" or "20" multiplier |  |  |  |  | Diode Linearity (Lin=(delta Pin (dB)/log(V1/V2)) | Video Resistance @ 23 Deg C | Power Level | Load resistor | Temperature (Degrees C): 23 deg data is the Output Voltage measured. At other Temperatures the data given is the drift relative to 23 deg in dB using the diode Linearity- not a simple "10" or "20" multiplier |  |  |  |  | Diode Linearity (Lin=(delta Pin (dB)/log(V1/V2))
dBm | K ohms | Drift @ -22 deg rel to 23 (dB) | Drift @ 0 deg rel to 23 (dB) | Vout @ 23 deg (V) | Drift @ 50 deg rel to 23 (dB) | Drift @ 71 deg rel to 23 (dB) | 10=Square law 20= Linear Region >20= Saturation neg = breakdown | Kohms | dBm | K ohms | Drift @ -22 deg rel to 23 (dB) | Drift @ 0 deg rel to 23 (dB) | Vout @ 23 deg (V) | Drift @ 50 deg rel to 23 (dB) | Drift @ 71 deg rel to 23 (dB) | 10=Square law 20= Linear Region >20= Saturation
18 | 10000 | 1.17 | 0.71 | 1.5924 | -0.791 | -1.856 | 26.28 |  | 18 | 10000 | 1.17 | 0.71 | 1.5924 | -0.791 | -1.856 | 26.28
18 | 200 | 1.17 | 0.71 | 1.5876 | -0.79 | -1.851 | 26.23 | 0.60469 | 15 | 10000 | 1.067 | 0.64 | 1.2243 | -0.83 | -1.35 | 23.96
18 | 50 | 1.157 | 0.7 | 1.5734 | -0.78 | -1.84 | 26.05 | 0.60379 | 13 | 10000 | 0.988 | 0.619 | 1.0103 | -0.723 | -1.294 | 23.3
18 | 20 | 1.127 | 0.68 | 1.546 | -0.77 | -1.806 | 25.76 | 0.60026 | 12 | 10000 | 0.931 | 0.605 | 0.91521 | -0.662 | -1.225 | 22.66
18 | 10 | 1.086 | 0.664 | 1.5017 | -0.74 | -1.75 | 25.32 | 0.60398 | 9 | 10000 | 0.681 | 0.428 | 0.67475 | -0.649 | -1.123 | 19.95
18 | 5 | 1.011 | 0.623 | 1.4204 | -0.7 | -1.68 | 24.69 | 0.60546 | 7 | 10000 | 0.722 | 0.44 | 0.53569 | -0.58 | -1.13 | 20.88
18 | 2 | 0.853 | 0.533 | 1.2247 | -0.62 | -1.54 | 23.45 | 0.60047 | 6 | 10000 | 0.553 | 0.303 | 0.47977 | -0.582 | -1.04 | 18.38
18 | 1 | 0.714 | 0.459 | 0.99885 | -0.54 | -1.4 | 22.35 | 0.59423 | 3 | 10000 | 0.512 | 0.348 | 0.32947 | -0.477 | -1.004 | 17.27
15 | 10000 | 1.067 | 0.64 | 1.2243 | -0.83 | -1.35 | 23.96 |  | 0 | 10000 | 0.588 | 0.358 | 0.22085 | -0.488 | -1.001 | 17.05
15 | 200 | 1.048 | 0.628 | 1.22 | -0.824 | -1.332 | 23.75 | 0.70492 | -3 | 10000 | 0.63395 | 0.34875 | 0.14729 | -0.49135 | -1.0214 | 15.5
15 | 50 | 1.019 | 0.617 | 1.207 | -0.818 | -1.306 | 23.55 | 0.71791 | -8 | 10000 | 0.73 | 0.36 | 0.071035 | -0.557 | -1.25 | 13.82
15 | 20 | 0.974 | 0.586 | 1.1822 | -0.8 | -1.28 | 23.42 | 0.71223 | -12 | 10000 | 0.739 | 0.477 | 0.036471 | -0.762 | -1.313 | 12.91
15 | 10 | 0.913 | 0.553 | 1.1431 | -0.75 | -1.222 | 23.15 | 0.71016 | -17 | 10000 | 1.007 | 0.623 | 0.01495 | -0.88 | -1.484 | 13.66
15 | 5 | 0.781 | 0.498 | 1.0737 | -0.699 | -1.137 | 22.46 | 0.7011 | -20 | 10000 | 1.108 | 0.604 | 0.009017 | -0.859 | -0.874 | 12.23
15 | 2 | 0.57 | 0.386 | 0.9122 | -0.601 | -0.955 | 21.25 | 0.68428 | [8 empty] | Rvideo @ 200K load
15 | 1 | 0.427 | 0.303 | 0.73328 | -0.507 | -0.797 | 20.41 | 0.66962 | 18 | 200 | 1.17 | 0.71 | 1.5876 | -0.79 | -1.851 | 26.23 | 0.60469
13 | 10000 | 0.988 | 0.619 | 1.0103 | -0.723 | -1.294 | 23.3 |  | 15 | 200 | 1.048 | 0.628 | 1.22 | -0.824 | -1.332 | 23.75 | 0
13 | 200 | 0.997 | 0.632 | 1.005 | -0.723 | -1.289 | 23.51 | 1.0587 | 13 | 200 | 0.997 | 0.632 | 1.005 | -0.723 | -1.289 | 23.51 | 1.0587
13 | 50 | 0.96 | 0.614 | 0.9926 | 0.698 | -1.26 | 23.47 | 0.89109 | 12 | 200 | 0.916 | 0.598 | 0.91119 | -0.651 | -1.229 | 22.62 | 0.88258
13 | 20 | 0.841 | 0.536 | 0.9711 | -0.665 | -1.205 | 22.65 | 0.80713 | 9 | 200 | 0.643 | 0.409 | 0.67144 | -0.634 | -1.11 | 19.84 | 0.98684
13 | 10 | 0.723 | 0.486 | 0.9369 | -0.612 | -1.098 | 21.74 | 0.78335 | 7 | 200 | 0.663 | 0.416 | 0.53238 | -0.564 | -1.107 | 20.79 | 1.2438
13 | 5 | 0.608 | 0.423 | 0.87465 | -0.551 | -0.995 | 21.24 | 0.7754 | 6 | 200 | 0.49 | 0.281 | 0.47656 | -0.57 | -1.01 | 18.29 | 1.3493
13 | 2 | 0.41 | 0.313 | 0.73449 | -0.44 | -0.806 | 20.31 | 0.751 | 3 | 200 | 0.409 | 0.31 | 0.32665 | -0.45 | -0.972 | 17.2 | 1.7217
13 | 1 | 0.27 | 0.232 | 0.58519 | -0.351 | -0.654 | 19.57 | 0.72643 | 0 | 200 | 0.405 | 0.295 | 0.21861 | -0.46 | -0.96 | 16.96 | 2.042
12 | 10000 | 0.931 | 0.605 | 0.91521 | -0.662 | -1.225 | 22.66 |  | -3 | 200 | 0.39116 | 0.27258 | 0.14549 | -0.45584 | -0.97174 | 15.4 | 2.4757
12 | 200 | 0.916 | 0.598 | 0.91119 | -0.651 | -1.229 | 22.62 | 0.88258 | -8 | 200 | 0.406 | 0.25 | 0.069979 | -0.52 | -1.194 | 13.78 | 3.018
12 | 50 | 0.851 | 0.572 | 0.8998 | -0.623 | -1.19 | 22.4 | 0.85636 | -12 | 200 | 0.369 | 0.355 | 0.035868 | -0.718 | -1.254 | 12.91 | 3.3623
12 | 20 | 0.776 | 0.54 | 0.87726 | -0.58 | -1.115 | 22.12 | 0.86524 | -17 | 200 | 0.508 | 0.482 | 0.014699 | -0.834 | -1.422 | 13.63 | 3.4152
12 | 10 | 0.701 | 0.499 | 0.84272 | -0.53 | -1.027 | 21.82 | 0.8602 | -20 | 200 | 0.663 | 0.479 | 0.008856 | -0.806 | -0.813 | 12.23 | 3.636
12 | 5 | 0.558 | 0.412 | 0.78479 | -0.46 | -0.912 | 21.15 | 0.83093 | [8 empty] | Rvideo @ 50 K load
12 | 2 | 0.352 | 0.297 | 0.65577 | -0.349 | -0.715 | 20.16 | 0.79125 | 18 | 50 | 1.157 | 0.7 | 1.5734 | -0.78 | -1.84 | 26.05 | 0.60379
12 | 1 | 0.218 | 0.218 | 0.52023 | -0.277 | -0.564 | 19.48 | 0.75925 | 15 | 50 | 1.019 | 0.617 | 1.207 | -0.818 | -1.306 | 23.55 | 0.53978
9 | 10000 | 0.681 | 0.428 | 0.67475 | -0.649 | -1.123 | 19.95 |  | 13 | 50 | 0.96 | 0.614 | 0.9926 | -0.698 | -1.26 | 23.47 | 0.89109
9 | 200 | 0.643 | 0.409 | 0.67144 | -0.634 | -1.11 | 19.84 | 0.98684 | 12 | 50 | 0.851 | 0.572 | 0.8998 | -0.623 | -1.19 | 22.4 | 0.85636
9 | 50 | 0.56 | 0.377 | 0.66109 | -0.604 | -1.066 | 19.68 | 1.0328 | 9 | 50 | 0.56 | 0.377 | 0.66109 | -0.604 | -1.066 | 19.68 | 1.0328
9 | 20 | 0.446 | 0.321 | 0.64192 | -0.554 | -0.991 | 19.34 | 1.0228 | 7 | 50 | 0.53 | 0.354 | 0.52314 | -0.53 | -1.05 | 20.52 | 1.1996
9 | 10 | 0.328 | 0.259 | 0.614 | -0.5 | -0.9 | 18.92 | 0.98938 | 6 | 50 | 0.352 | 0.216 | 0.46762 | -0.53 | -0.955 | 18.06 | 1.2992
9 | 5 | 0.198 | 0.185 | 0.56611 | -0.428 | -0.771 | 18.42 | 0.95948 | 3 | 50 | 0.204 | 0.217 | 0.31902 | -0.4 | -0.89 | 16.98 | 1.6366
9 | 2 | 0.045 | 0.094 | 0.46551 | -0.313 | -0.572 | 17.68 | 0.89896 | 0 | 50 | 0.088 | 0.163 | 0.21241 | -0.39 | -0.85 | 16.75 | 1.9855
9 | 1 | -0.049 | 0.035 | 0.36491 | -0.233 | -0.432 | 17.15 | 0.84908 | -3 | 50 | -0.035075 | 0.1037 | 0.14064 | -0.3721 | -0.85095 | 15.25 | 2.3671
7 | 10000 | 0.722 | 0.44 | 0.53569 | -0.58 | -1.13 | 20.88 |  | -8 | 50 | -0.217 | -0.004 | 0.067093 | -0.409 | -1.045 | 13.68 | 2.9377
7 | 200 | 0.663 | 0.416 | 0.53238 | -0.564 | -1.107 | 20.79 | 1.2438 | -12 | 50 | -0.413 | 0.047 | 0.034218 | -0.602 | -1.092 | 12.85 | 3.2921
7 | 50 | 0.53 | 0.354 | 0.52314 | -0.53 | -1.05 | 20.52 | 1.1996 | -17 | 50 | -0.592 | 0.11 | 0.013965 | -0.696 | -1.233 | 13.63 | 3.5267
7 | 20 | 0.375 | 0.276 | 0.50587 | -0.466 | -0.955 | 20.19 | 1.179 | -20 | 50 | -0.366 | 0.143 | 0.008414 | -0.681 | -0.638 | 12.18 | 3.5833
7 | 10 | 0.235 | 0.195 | 0.48137 | -0.398 | -0.839 | 19.77 | 1.1285 | [8 empty] | Rvideo @ 20 K load
7 | 5 | 0.096 | 0.104 | 0.44091 | -0.31 | -0.69 | 19.21 | 1.0748 | 18 | 20 | 1.127 | 0.68 | 1.546 | -0.77 | -1.806 | 25.76 | 0.60026
7 | 2 | -0.052 | 0.005 | 0.35878 | -0.19 | -0.47 | 18.51 | 0.98616 | 15 | 20 | 0.974 | 0.586 | 1.1822 | -0.8 | -1.28 | 23.42 | 0.63949
7 | 1 | -0.137 | -0.054 | 0.27896 | -0.105 | -0.317 | 18.09 | 0.92031 | 13 | 20 | 0.841 | 0.536 | 0.9711 | -0.665 | -1.205 | 22.65 | 0.80713
6 | 10000 | 0.553 | 0.303 | 0.47977 | -0.582 | -1.04 | 18.38 |  | 12 | 20 | 0.776 | 0.54 | 0.87726 | -0.58 | -1.115 | 22.12 | 0.86524
6 | 200 | 0.49 | 0.281 | 0.47656 | -0.57 | -1.01 | 18.29 | 1.3493 | 9 | 20 | 0.446 | 0.321 | 0.64192 | -0.554 | -0.991 | 19.34 | 1.0228
6 | 50 | 0.352 | 0.216 | 0.46762 | -0.53 | -0.955 | 18.06 | 1.2992 | 7 | 20 | 0.375 | 0.276 | 0.50587 | -0.466 | -0.955 | 20.19 | 1.179
6 | 20 | 0.2 | 0.133 | 0.45135 | -0.47 | -0.86 | 17.74 | 1.2591 | 6 | 20 | 0.2 | 0.133 | 0.45135 | -0.47 | -0.86 | 17.74 | 1.2591
6 | 10 | 0.069 | 0.055 | 0.42846 | -0.399 | -0.747 | 17.34 | 1.1975 | 3 | 20 | -0.002 | 0.104 | 0.30577 | -0.323 | -0.765 | 16.64 | 1.5499
6 | 5 | -0.053 | -0.026 | 0.39111 | -0.315 | -0.603 | 16.87 | 1.1334 | 0 | 20 | -0.213 | 0.003 | 0.20187 | -0.277 | -0.687 | 16.39 | 1.88
6 | 2 | -0.18 | -0.117 | 0.31682 | -0.21 | -0.401 | 16.17 | 1.0287 | -3 | 20 | -0.43806 | -0.09536 | 0.13245 | -0.23989 | -0.6556 | 14.9 | 2.2415
6 | 1 | -0.253 | -0.169 | 0.24563 | -0.14 | -0.27 | 15.71 | 0.95322 | -8 | 20 | -0.875 | -0.327 | 0.062236 | -0.239 | -0.798 | 13.48 | 2.8276
3 | 10000 | 0.512 | 0.348 | 0.32947 | -0.477 | -1.004 | 17.27 |  | -12 | 20 | -1.335 | -0.374 | 0.031423 | -0.406 | -0.813 | 12.74 | 3.2129
3 | 200 | 0.409 | 0.31 | 0.32665 | -0.45 | -0.972 | 17.2 | 1.7217 | -17 | 20 | -1.978 | -0.44 | 0.012728 | -0.45 | -0.9 | 13.57 | 3.4915
3 | 50 | 0.204 | 0.217 | 0.31902 | -0.4 | -0.89 | 16.98 | 1.6366 | -20 | 20 | -1.694 | -0.358 | 0.007649 | -0.456 | -0.336 | 12.14 | 3.5769
3 | 20 | -0.002 | 0.104 | 0.30577 | -0.323 | -0.765 | 16.64 | 1.5499 | [8 empty] | Rvideo @10 K load
3 | 10 | -0.154 | 0.01 | 0.2877 | -0.236 | -0.626 | 16.26 | 1.4518 | 18 | 10 | 1.086 | 0.664 | 1.5017 | -0.74 | -1.75 | 25.32 | 0.60398
3 | 5 | -0.286 | -0.082 | 0.25969 | -0.14 | -0.45 | 15.74 | 1.3434 | 15 | 10 | 0.913 | 0.553 | 1.1431 | -0.75 | -1.222 | 23.15 | 0.67255
3 | 2 | -0.409 | -0.17 | 0.20667 | -0.016 | -0.229 | 15.03 | 1.1883 | 13 | 10 | 0.723 | 0.486 | 0.9369 | -0.612 | -1.098 | 21.74 | 0.78335
3 | 1 | -0.472 | -0.22 | 0.15822 | 0.052 | -0.091 | 14.51 | 1.0823 | 12 | 10 | 0.701 | 0.499 | 0.84272 | -0.53 | -1.027 | 21.82 | 0.8602
0 | 10000 | 0.588 | 0.358 | 0.22085 | -0.488 | -1.001 | 17.05 |  | 9 | 10 | 0.328 | 0.259 | 0.614 | -0.5 | -0.9 | 18.92 | 0.98938
0 | 200 | 0.405 | 0.295 | 0.21861 | -0.46 | -0.96 | 16.96 | 2.042 | 7 | 10 | 0.235 | 0.195 | 0.48137 | -0.398 | -0.839 | 19.77 | 1.1285
0 | 50 | 0.088 | 0.163 | 0.21241 | -0.39 | -0.85 | 16.75 | 1.9855 | 6 | 10 | 0.069 | 0.055 | 0.42846 | -0.399 | -0.747 | 17.34 | 1.1975
0 | 20 | -0.213 | 0.003 | 0.20187 | -0.277 | -0.687 | 16.39 | 1.88 | 3 | 10 | -0.154 | 0.01 | 0.2877 | -0.236 | -0.626 | 16.26 | 1.4518
0 | 10 | -0.423 | -0.122 | 0.18812 | -0.163 | -0.503 | 15.94 | 1.7395 | 0 | 10 | -0.423 | -0.122 | 0.18812 | -0.163 | -0.503 | 15.94 | 1.7395
0 | 5 | -0.596 | -0.237 | 0.16746 | -0.03 | -0.279 | 15.34 | 1.5941 | -3 | 10 | -0.7279 | -0.261 | 0.12196 | -0.0986 | -0.4321 | 14.5 | 2.0771
0 | 2 | -0.744 | -0.343 | 0.13054 | 0.121 | -0.002 | 14.39 | 1.3836 | -8 | 10 | -1.413 | -0.623 | 0.052609 | -0.041 | -0.499 | 13.17 | 3.5024
0 | 1 | -0.811 | -0.393 | 0.098278 | 0.201 | 0.16 | 13.68 | 1.2471 | -12 | 10 | -2.178 | -0.795 | 0.027873 | -0.162 | -0.462 | 12.55 | 3.0847
-3 | 10000 | 0.63395 | 0.34875 | 0.14729 | -0.49135 | -1.0214 | 15.5 |  | -17 | 10 | -3.289 | -1.029 | 0.011139 | -0.133 | -0.449 | 13.45 | 3.4213
-3 | 200 | 0.39116 | 0.27258 | 0.14549 | -0.45584 | -0.97174 | 15.4 | 2.4757 | -20 | 10 | -2.958 | -0.913 | 0.006664 | -0.16 | 0.078 | 12.06 | 3.5309
-3 | 50 | -0.035075 | 0.1037 | 0.14064 | -0.3721 | -0.85095 | 15.25 | 2.3671 | [8 empty] | Rvideo @ 5K load
-3 | 20 | -0.43806 | -0.09536 | 0.13245 | -0.23989 | -0.6556 | 14.9 | 2.2415 | 18 | 5 | 1.011 | 0.623 | 1.4204 | -0.7 | -1.68 | 24.69 | 0.60546
-3 | 10 | -0.7279 | -0.261 | 0.12196 | -0.0986 | -0.4321 | 14.5 | 2.0771 | 15 | 5 | 0.781 | 0.498 | 1.0737 | -0.699 | -1.137 | 22.46 | 0.68108
-3 | 5 | -0.96883 | -0.41283 | 0.10673 | 0.06255 | -0.16541 | 13.9 | 1.8999 | 13 | 5 | 0.608 | 0.423 | 0.87465 | -0.551 | -0.995 | 21.24 | 0.7754
-3 | 2 | -1.1739 | -0.55212 | 0.08077 | 0.24897 | 0.17157 | 12.9 | 1.6472 | 12 | 5 | 0.558 | 0.412 | 0.78479 | -0.46 | -0.912 | 21.15 | 0.83093
-3 | 1 | -1.2725 | -0.61854 | 0.059319 | 0.34892 | 0.36722 | 12.2 | 1.4831 | 9 | 5 | 0.198 | 0.185 | 0.56611 | -0.428 | -0.771 | 18.42 | 0.95948
-8 | 10000 | 0.73 | 0.36 | 0.071035 | -0.557 | -1.25 | 13.82 |  | 7 | 5 | 0.096 | 0.104 | 0.44091 | -0.31 | -0.69 | 19.21 | 1.0748
-8 | 200 | 0.406 | 0.25 | 0.069979 | -0.52 | -1.194 | 13.78 | 3.018 | 6 | 5 | -0.053 | -0.026 | 0.39111 | -0.315 | -0.603 | 16.87 | 1.1334
-8 | 50 | -0.217 | -0.004 | 0.067093 | -0.409 | -1.045 | 13.68 | 2.9377 | 3 | 5 | -0.286 | -0.082 | 0.25969 | -0.14 | -0.45 | 15.74 | 1.3434
-8 | 20 | -0.875 | -0.327 | 0.062236 | -0.239 | -0.798 | 13.48 | 2.8276 | 0 | 5 | -0.596 | -0.237 | 0.16746 | -0.03 | -0.279 | 15.34 | 1.5941
-8 | 10 | -1.413 | -0.623 | 0.052609 | -0.041 | -0.499 | 13.17 | 3.5024 | -3 | 5 | -0.96883 | -0.41283 | 0.10673 | 0.06255 | -0.16541 | 13.9 | 1.8999
-8 | 5 | -1.92 | -0.923 | 0.047433 | 0.198 | -0.113 | 12.66 | 2.4879 | -8 | 5 | -1.92 | -0.923 | 0.047433 | 0.198 | -0.113 | 12.66 | 2.4879
-8 | 2 | -2.429 | -1.236 | 0.033655 | 0.501 | 0.407 | 11.74 | 2.2214 | -12 | 5 | -3.048 | -1.26 | 0.022915 | 0.163 | 0.026 | 12.21 | 2.9579
-8 | 1 | -2.675 | -1.392 | 0.023347 | 0.674 | 0.729 | 11.02 | 2.0426 | -17 | 5 | -4.672 | -1.714 | 0.008927 | 0.314 | 0.203 | 13.23 | 3.3735
-12 | 10000 | 0.739 | 0.477 | 0.036471 | -0.762 | -1.313 | 12.91 |  | -20 | 5 | -4.289 | -1.579 | 0.005296 | 0.279 | 0.71 | 11.96 | 3.513
-12 | 200 | 0.369 | 0.355 | 0.035868 | -0.718 | -1.254 | 12.91 | 3.3623 | [8 empty] | Rvideo @ 2 K load
-12 | 50 | -0.413 | 0.047 | 0.034218 | -0.602 | -1.092 | 12.85 | 3.2921 | 18 | 2 | 0.853 | 0.533 | 1.2247 | -0.62 | -1.54 | 23.45 | 0.60047
-12 | 20 | -1.335 | -0.374 | 0.031423 | -0.406 | -0.813 | 12.74 | 3.2129 | 15 | 2 | 0.57 | 0.386 | 0.9122 | -0.601 | -0.955 | 21.25 | 0.67485
-12 | 10 | -2.178 | -0.795 | 0.027873 | -0.162 | -0.462 | 12.55 | 3.0847 | 13 | 2 | 0.41 | 0.313 | 0.73449 | -0.44 | -0.806 | 20.31 | 0.751
-12 | 5 | -3.048 | -1.26 | 0.022915 | 0.163 | 0.026 | 12.21 | 2.9579 | 12 | 2 | 0.352 | 0.297 | 0.65577 | -0.349 | -0.715 | 20.16 | 0.79125
-12 | 2 | -3.995 | -1.798 | 0.015363 | 0.613 | 0.751 | 11.59 | 2.7479 | 9 | 2 | 0.045 | 0.094 | 0.46551 | -0.313 | -0.572 | 17.68 | 0.89896
-12 | 1 | -4.488 | -2.098 | 0.010123 | 0.899 | 1.235 | 11.07 | 2.6028 | 7 | 2 | -0.052 | 0.005 | 0.35878 | -0.19 | -0.47 | 18.51 | 0.98616
-17 | 10000 | 1.007 | 0.623 | 0.01495 | -0.88 | -1.484 | 13.66 |  | 6 | 2 | -0.18 | -0.117 | 0.31682 | -0.21 | -0.401 | 16.17 | 1.0287
-17 | 200 | 0.508 | 0.482 | 0.014699 | 0.834 | -1.422 | 13.63 | 3.4152 | 3 | 2 | -0.409 | -0.17 | 0.20667 | -0.016 | -0.229 | 15.03 | 1.1883
-17 | 50 | -0.592 | 0.11 | 0.013965 | -0.696 | -1.233 | 13.63 | 3.5267 | 0 | 2 | -0.744 | -0.343 | 0.13054 | 0.121 | -0.002 | 14.39 | 1.3836
-17 | 20 | -1.978 | -0.44 | 0.012728 | -0.45 | -0.9 | 13.57 | 3.4915 | -3 | 2 | -1.1739 | -0.55212 | 0.08077 | 0.24897 | 0.17157 | 12.9 | 1.6472
-17 | 10 | -3.289 | -1.029 | 0.011139 | -0.133 | -0.449 | 13.45 | 3.4213 | -8 | 2 | -2.429 | -1.236 | 0.033655 | 0.501 | 0.407 | 11.74 | 2.2214
-17 | 5 | -4.672 | -1.714 | 0.008927 | 0.314 | 0.203 | 13.23 | 3.3735 | -12 | 2 | -3.995 | -1.798 | 0.015363 | 0.613 | 0.751 | 11.59 | 2.7479
-17 | 2 | -6.131 | -2.546 | 0.005688 | 0.98 | 1.229 | 12.79 | 3.2567 | -17 | 2 | -6.131 | -2.546 | 0.005688 | 0.98 | 1.229 | 12.79 | 3.2567
-17 | 1 | -6.906 | -2.974 | 0.003577 | 1.439 | 1.973 | 12.5 | 3.1795 | -20 | 2 | -5.732 | -2.374 | 0.003315 | 0.953 | 1.723 | 11.7 | 3.4401
-20 | 10000 | 1.108 | 0.604 | 0.009017 | -0.859 | -0.874 | 12.23 | [9 empty] | Rvideo @ 1 K load
-20 | 200 | 0.663 | 0.479 | 0.008856 | -0.806 | -0.813 | 12.23 | 3.636 | 18 | 1 | 0.714 | 0.459 | 0.99885 | -0.54 | -1.4 | 22.35 | 0.59423
-20 | 50 | -0.366 | 0.143 | 0.008414 | -0.681 | -0.638 | 12.18 | 3.5833 | 15 | 1 | 0.427 | 0.303 | 0.73328 | -0.507 | -0.797 | 20.41 | 0.66376
-20 | 20 | -1.694 | -0.358 | 0.007649 | -0.456 | -0.336 | 12.14 | 3.5769 | 13 | 1 | 0.27 | 0.232 | 0.58519 | -0.351 | -0.654 | 19.57 | 0.72643
-20 | 10 | -2.958 | -0.913 | 0.006664 | -0.16 | 0.078 | 12.06 | 3.5309 | 12 | 1 | 0.218 | 0.218 | 0.52023 | -0.277 | -0.564 | 19.48 | 0.75925
-20 | 5 | -4.289 | -1.579 | 0.005296 | 0.279 | 0.71 | 11.96 | 3.513 | 9 | 1 | -0.049 | 0.035 | 0.36491 | -0.233 | -0.432 | 17.15 | 0.84908
-20 | 2 | -5.732 | -2.374 | 0.003315 | 0.953 | 1.723 | 11.7 | 3.4401 | 7 | 1 | -0.137 | -0.054 | 0.27896 | -0.105 | -0.317 | 18.09 | 0.92031
-20 | 1 | -6.446 | -2.827 | 0.002058 | 1.385 | 2.42 | 11.4 | 3.3814 | 6 | 1 | -0.253 | -0.169 | 0.24563 | -0.14 | -0.27 | 15.71 | 0.95322
[9 empty] | 3 | 1 | -0.472 | -0.22 | 0.15822 | 0.052 | -0.091 | 14.51 | 1.0823
Test Temperatures | [8 empty] | 0 | 1 | -0.811 | -0.393 | 0.098278 | 0.201 | 0.16 | 13.68 | 1.2471
-22 | Extended low- non-destructive but electrically not specified | [7 empty] | -3 | 1 | -1.2725 | -0.61854 | 0.059319 | 0.34892 | 0.36722 | 12.2 | 1.4831
0 | Lower Temperature Spec | [7 empty] | -8 | 1 | -2.675 | -1.392 | 0.023347 | 0.674 | 0.729 | 11.02 | 2.0426
23 | Room Temp | [7 empty] | -12 | 1 | -4.488 | -2.098 | 0.010123 | 0.899 | 1.235 | 11.07 | 2.6028
50 | Upper Temperature Spec | [7 empty] | -17 | 1 | -6.906 | -2.974 | 0.003577 | 1.439 | 1.973 | 12.5 | 3.1795
71 | Extended High- non-destructive but electrically not specified | [7 empty] | -20 | 1 | -6.446 | -2.827 | 0.002058 | 1.385 | 2.42 | 11.4 | 3.3814
[9 empty] | Power Level | Load resistor | Temperature (Degrees C): Average detector voltage output at the five test temperatures. Grouped by load resistance. |  |  |  |  | Diode Linearity (Lin=(delta Pin (dB)/log(V1/V2))
[9 empty] | dBm | K ohms | Vout @ -22 deg (V) | Vout @ 0 deg (V) | Vout @ 23 deg (V) | Vout @ 50 deg (V) | Vout @ 71 deg (V) | 10=Square law 20= Linear Region >20= Saturation
[9 empty] | 18 | 10000 | 1.7643 | 1.6946 | 1.5924 | 1.4858 | 1.3534 | 26.28
[9 empty] | 15 | 10000 | 1.3565 | 1.302 | 1.2243 | 1.1304 | 1.0753 | 23.96
[9 empty] | 13 | 10000 | 1.1139 | 1.074 | 1.0103 | 0.94062 | 0.88902 | 23.3
[9 empty] | 12 | 10000 | 1.006 | 0.97324 | 0.91521 | 0.85567 | 0.80809 | 22.66
[9 empty] | 9 | 10000 | 0.72992 | 0.70892 | 0.67475 | 0.62605 | 0.59272 | 19.95
[9 empty] | 7 | 10000 | 0.58009 | 0.56232 | 0.53569 | 0.5025 | 0.47293 | 20.88
[9 empty] | 6 | 10000 | 0.51419 | 0.49833 | 0.47977 | 0.44603 | 0.42116 | 18.38
[9 empty] | 3 | 10000 | 0.35274 | 0.34511 | 0.32947 | 0.30917 | 0.28819 | 17.27
[9 empty] | 0 | 10000 | 0.2391 | 0.23178 | 0.22085 | 0.20676 | 0.19292 | 17.05
[9 empty] | -3 | 10000 | 0.16184 | 0.15513 | 0.14729 | 0.13692 | 0.12656 | 15.5
[9 empty] | -8 | 10000 | 0.080222 | 0.075426 | 0.071035 | 0.064739 | 0.05768 | 13.82
[9 empty] | -12 | 10000 | 0.041609 | 0.03971 | 0.036471 | 0.031836 | 0.028856 | 12.91
[9 empty] | -17 | 10000 | 0.017716 | 0.016605 | 0.01495 | 0.012889 | 0.011641 | 13.66
[9 empty] | -20 | 10000 | 0.011109 | 0.010103 | 0.009017 | 0.0076705 | 0.0076489 | 12.23
[9 empty] | 18 | 200 | 1.7593 | 1.6897 | 1.5876 | 1.4812 | 1.3495 | 26.23
[9 empty] | 15 | 200 | 1.3505 | 1.2966 | 1.22 | 1.1263 | 1.0722 | 23.75
[9 empty] | 13 | 200 | 1.1081 | 1.0691 | 1.005 | 0.93627 | 0.88578 | 23.51
[9 empty] | 12 | 200 | 1.0002 | 0.96838 | 0.91119 | 0.85276 | 0.80404 | 22.62
[9 empty] | 9 | 200 | 0.72346 | 0.70407 | 0.67144 | 0.6238 | 0.59028 | 19.84
[9 empty] | 7 | 200 | 0.57294 | 0.55748 | 0.53238 | 0.50014 | 0.47095 | 20.79
[9 empty] | 6 | 200 | 0.50688 | 0.49372 | 0.47656 | 0.44356 | 0.41965 | 18.29
[9 empty] | 3 | 200 | 0.34504 | 0.3405 | 0.32665 | 0.30756 | 0.2868 | 17.2
[9 empty] | 0 | 200 | 0.23097 | 0.22755 | 0.21861 | 0.20538 | 0.1919 | 16.96
[9 empty] | -3 | 200 | 0.15425 | 0.15154 | 0.14549 | 0.13591 | 0.12582 | 15.4
[9 empty] | -8 | 200 | 0.074891 | 0.072964 | 0.069979 | 0.064155 | 0.057322 | 13.78
[9 empty] | -12 | 200 | 0.038308 | 0.038212 | 0.035868 | 0.031557 | 0.02868 | 12.91
[9 empty] | -17 | 200 | 0.016016 | 0.015946 | 0.014699 | 0.012767 | 0.01156 | 13.63
[9 empty] | -20 | 200 | 0.010033 | 0.0096918 | 0.008856 | 0.0076091 | 0.0075991 | 12.23
[9 empty] | 18 | 50 | 1.7428 | 1.6738 | 1.5734 | 1.4686 | 1.3372 | 26.05
[9 empty] | 15 | 50 | 1.3334 | 1.282 | 1.207 | 1.1142 | 1.0623 | 23.55
[9 empty] | 13 | 50 | 1.0906 | 1.0542 | 0.9926 | 0.9269 | 0.87718 | 23.47
[9 empty] | 12 | 50 | 0.98206 | 0.95429 | 0.8998 | 0.84398 | 0.7962 | 22.4
[9 empty] | 9 | 50 | 0.70586 | 0.69091 | 0.66109 | 0.61599 | 0.58357 | 19.68
[9 empty] | 7 | 50 | 0.5552 | 0.54434 | 0.52314 | 0.49293 | 0.465 | 20.52
[9 empty] | 6 | 50 | 0.48908 | 0.48068 | 0.46762 | 0.43706 | 0.41401 | 18.06
[9 empty] | 3 | 50 | 0.32797 | 0.32855 | 0.31902 | 0.30218 | 0.28275 | 16.98
[9 empty] | 0 | 50 | 0.215 | 0.21722 | 0.21241 | 0.20132 | 0.18899 | 16.75
[9 empty] | -3 | 50 | 0.13989 | 0.14285 | 0.14064 | 0.13295 | 0.12368 | 15.25
[9 empty] | -8 | 50 | 0.064687 | 0.067048 | 0.067093 | 0.06263 | 0.056271 | 13.68
[9 empty] | -12 | 50 | 0.031777 | 0.034507 | 0.034218 | 0.030719 | 0.028137 | 12.85
[9 empty] | -17 | 50 | 0.012636 | 0.014227 | 0.013965 | 0.012416 | 0.011339 | 13.63
[9 empty] | -20 | 50 | 0.0078515 | 0.0086446 | 0.008414 | 0.0073976 | 0.007458 | 12.18
[11 empty] | Vout @ -22 deg (V) | Vout @ 0 deg (V) | Vout @ 23 deg (V) | Vout @ 50 deg (V) | Vout @ 71 deg (V) | Diode Linearity | Rv @ 20K load, -22 C | Rv @ 20K load, 0 C | Rv @ 20K load, 23 C | Rv @ 20K load, 50 C | Rv @ 20K load,71 C
[9 empty] | 18 | 20 | 1.7099 | 1.6429 | 1.546 | 1.4432 | 1.3155 | 25.76 | 0.63684 | 0.62965 | 0.60026 | 0.59041 | 0.57583
[9 empty] | 15 | 20 | 1.301 | 1.2523 | 1.1822 | 1.0928 | 1.0424 | 23.42 | 0.85312 | 0.79298 | 0.71223 | 0.68925 | 0.6318
[9 empty] | 13 | 20 | 1.0578 | 1.0255 | 0.9711 | 0.90762 | 0.85914 | 22.65 | 1.0613 | 0.94663 | 0.80713 | 0.72726 | 0.69553
[9 empty] | 12 | 20 | 0.95106 | 0.92798 | 0.87726 | 0.82586 | 0.78113 | 22.12 | 1.1557 | 0.97539 | 0.86524 | 0.7219 | 0.69046
[9 empty] | 9 | 20 | 0.67693 | 0.66693 | 0.64192 | 0.60095 | 0.57048 | 19.34 | 1.5658 | 1.2592 | 1.0228 | 0.83554 | 0.77978
[9 empty] | 7 | 20 | 0.52797 | 0.52205 | 0.50587 | 0.47969 | 0.45367 | 20.19 | 1.9741 | 1.5431 | 1.179 | 0.95117 | 0.8491
[9 empty] | 6 | 20 | 0.46322 | 0.45921 | 0.45135 | 0.42464 | 0.40368 | 17.74 | 2.2003 | 1.7037 | 1.2591 | 1.0075 | 0.86607
[9 empty] | 3 | 20 | 0.30569 | 0.3102 | 0.30577 | 0.2924 | 0.27506 | 16.64 | 3.0788 | 2.2508 | 1.5499 | 1.1464 | 0.95481
[9 empty] | 0 | 20 | 0.19592 | 0.20195 | 0.20187 | 0.19416 | 0.1833 | 16.39 | 4.4079 | 2.9542 | 1.88 | 1.2974 | 1.0501
[9 empty] | -3 | 20 | 0.12378 | 0.13051 | 0.13245 | 0.12763 | 0.11969 | 14.9 | 6.1495 | 3.7719 | 2.2415 | 1.4567 | 1.1477
[9 empty] | -8 | 20 | 0.053596 | 0.058855 | 0.062236 | 0.059746 | 0.054305 | 13.48 | 9.936 | 5.6311 | 2.8276 | 1.6714 | 1.2428
[9 empty] | -12 | 20 | 0.024687 | 0.029369 | 0.031423 | 0.0292 | 0.027129 | 12.74 | 13.7101 | 7.0417 | 3.2129 | 1.8059 | 1.2736
[9 empty] | -17 | 20 | 0.0090991 | 0.011812 | 0.012728 | 0.011792 | 0.010925 | 13.57 | 18.9397 | 8.1153 | 3.4915 | 1.86 | 1.3106
[9 empty] | -20 | 20 | 0.0055471 | 0.0071469 | 0.007649 | 0.0070152 | 0.0071767 | 12.14 | 20.0519 | 8.2725 | 3.5769 | 1.8682 | 1.3158
[17 empty] | Rv @ 10K load, -22 C | Rv @ 10K load, 0 C | Rv @ 10K load, 23 C | Rv @ 10K load, 50 C | Rv @ 10K load,71 C
[9 empty] | 18 | 10 | 1.6576 | 1.5952 | 1.5017 | 1.404 | 1.2808 | 25.32 | 0.64384 | 0.62335 | 0.60398 | 0.58269 | 0.56721
[9 empty] | 15 | 10 | 1.2518 | 1.2078 | 1.1431 | 1.0609 | 1.0123 | 23.15 | 0.83653 | 0.78002 | 0.71016 | 0.65497 | 0.62278
[9 empty] | 13 | 10 | 1.0115 | 0.98639 | 0.9369 | 0.87809 | 0.83404 | 21.74 | 1.0129 | 0.88843 | 0.78335 | 0.7121 | 0.65917
[9 empty] | 12 | 10 | 0.90742 | 0.88828 | 0.84272 | 0.79688 | 0.75616 | 21.82 | 1.0866 | 0.9564 | 0.8602 | 0.73775 | 0.68673
[9 empty] | 9 | 10 | 0.63901 | 0.63366 | 0.614 | 0.57775 | 0.5503 | 18.92 | 1.4228 | 1.1876 | 0.98938 | 0.836 | 0.77088
[9 empty] | 7 | 10 | 0.49473 | 0.49243 | 0.48137 | 0.45957 | 0.43656 | 19.77 | 1.7254 | 1.4194 | 1.1285 | 0.93425 | 0.83315
[9 empty] | 6 | 10 | 0.43241 | 0.4316 | 0.42846 | 0.40635 | 0.388 | 17.34 | 1.8913 | 1.5461 | 1.1975 | 0.97656 | 0.85468
[9 empty] | 3 | 10 | 0.28149 | 0.28811 | 0.2877 | 0.27824 | 0.26329 | 16.26 | 2.5312 | 1.9787 | 1.4518 | 1.1114 | 0.94555
[9 empty] | 0 | 10 | 0.17697 | 0.18484 | 0.18812 | 0.18374 | 0.17494 | 15.94 | 3.5105 | 2.54 | 1.7395 | 1.2526 | 1.0279
[9 empty] | -3 | 10 | 0.10865 | 0.11701 | 0.12196 | 0.12007 | 0.11387 | 14.5 | 4.8957 | 3.2575 | 2.0771 | 1.4041 | 1.1138
[9 empty] | -8 | 10 | 0.041093 | 0.04718 | 0.052609 | 0.052233 | 0.048214 | 13.17 | 9.5219 | 5.9869 | 3.5024 | 2.3943 | 1.9633
[9 empty] | -12 | 10 | 0.018691 | 0.02409 | 0.027873 | 0.027057 | 0.025608 | 12.55 | 12.2614 | 6.4839 | 3.0847 | 1.7665 | 1.2687
[9 empty] | -17 | 10 | 0.0063432 | 0.0093399 | 0.011139 | 0.010888 | 0.010315 | 13.45 | 17.9286 | 7.779 | 3.4213 | 1.8375 | 1.286
[9 empty] | -20 | 10 | 0.0037884 | 0.005598 | 0.006664 | 0.0064635 | 0.006764 | 12.06 | 19.3223 | 8.0476 | 3.5309 | 1.8674 | 1.3083
[11 empty] | Vout @ -22 deg (V) | Vout @ 0 deg (V) | Vout @ 23 deg (V) | Vout @ 50 deg (V) | Vout @ 71 deg (V) | Diode Linearity @ 5 K load | Rv @ 5K load, -22 C | Rv @ 5K load, 0 C | Rv @ 5K load, 23 C | Rv @ 5K load, 50 C | Rv @ 5K load,71 C
[9 empty] | 18 | 5 | 1.5608 | 1.5054 | 1.4204 | 1.3306 | 1.2144 | 24.69 | 0.65177 | 0.62853 | 0.60546 | 0.58296 | 0.57226
[9 empty] | 15 | 5 | 1.1632 | 1.13 | 1.0737 | 0.99949 | 0.9556 | 22.46 | 0.83067 | 0.76099 | 0.68108 | 0.65509 | 0.6265
[9 empty] | 13 | 5 | 0.93424 | 0.91569 | 0.87465 | 0.82393 | 0.78521 | 21.24 | 0.96157 | 0.86453 | 0.7754 | 0.70813 | 0.66098
[9 empty] | 12 | 5 | 0.83394 | 0.82079 | 0.78479 | 0.74646 | 0.71061 | 21.15 | 1.0317 | 0.92866 | 0.83093 | 0.73156 | 0.68589
[9 empty] | 9 | 5 | 0.5803 | 0.57936 | 0.56611 | 0.53662 | 0.5141 | 18.42 | 1.2892 | 1.1181 | 0.95948 | 0.83328 | 0.76467
[9 empty] | 7 | 5 | 0.44601 | 0.44644 | 0.44091 | 0.42483 | 0.40591 | 19.21 | 1.503 | 1.2979 | 1.0748 | 0.91417 | 0.82552
[9 empty] | 6 | 5 | 0.38829 | 0.38972 | 0.39111 | 0.37465 | 0.36021 | 16.87 | 1.6211 | 1.3934 | 1.1334 | 0.95267 | 0.84608
[9 empty] | 3 | 5 | 0.24905 | 0.25659 | 0.25969 | 0.25443 | 0.24315 | 15.74 | 2.0818 | 1.7249 | 1.3434 | 1.0757 | 0.92624
[9 empty] | 0 | 5 | 0.15313 | 0.1616 | 0.16746 | 0.1667 | 0.16059 | 15.34 | 2.8072 | 2.1714 | 1.5941 | 1.2014 | 1.0067
[9 empty] | -3 | 5 | 0.090909 | 0.09968 | 0.10673 | 0.10785 | 0.10385 | 13.9 | 3.9011 | 2.7812 | 1.8999 | 1.3481 | 1.0932
[9 empty] | -8 | 5 | 0.033452 | 0.040103 | 0.047433 | 0.049172 | 0.046468 | 12.66 | 6.9906 | 4.4041 | 2.4879 | 1.5829 | 1.2064
[9 empty] | -12 | 5 | 0.012897 | 0.018069 | 0.022915 | 0.02363 | 0.023028 | 12.21 | 11.1314 | 5.9885 | 2.9579 | 1.7363 | 1.2656
[9 empty] | -17 | 5 | 0.0039588 | 0.0066245 | 0.008927 | 0.0094284 | 0.009248 | 13.23 | 17.3749 | 7.5334 | 3.3735 | 1.8352 | 1.294
[9 empty] | -20 | 5 | 0.0023192 | 0.0039077 | 0.005296 | 0.0055882 | 0.0060717 | 11.96 | 18.9493 | 7.9269 | 3.513 | 1.8631 | 1.2988
[11 empty] | Vout @ -22 deg (V) | Vout @ 0 deg (V) | Vout @ 23 deg (V) | Vout @ 50 deg (V) | Vout @ 71 deg (V) |  | Rv @ 2K load, -22 C | Rv @ 2K load, 0 C | Rv @ 2K load, 23 C | Rv @ 2K load, 50 C | Rv @ 2K load,71 C
[9 empty] | 18 | 2 | 1.3317 | 1.2905 | 1.2247 | 1.1524 | 1.0528 | 23.45 | 0.64971 | 0.62627 | 0.60047 | 0.57865 | 0.57099
[9 empty] | 15 | 2 | 0.97032 | 0.95116 | 0.9122 | 0.85469 | 0.82252 | 21.25 | 0.796 | 0.73763 | 0.68428 | 0.64526 | 0.61472
[9 empty] | 13 | 2 | 0.76944 | 0.76102 | 0.73449 | 0.69875 | 0.67035 | 20.31 | 0.89539 | 0.82258 | 0.751 | 0.69231 | 0.6524
[9 empty] | 12 | 2 | 0.68267 | 0.6784 | 0.65577 | 0.63015 | 0.60435 | 20.16 | 0.9473 | 0.86923 | 0.79125 | 0.71579 | 0.67427
[9 empty] | 9 | 2 | 0.46825 | 0.47124 | 0.46551 | 0.44692 | 0.43209 | 17.68 | 1.1177 | 1.0087 | 0.89896 | 0.80166 | 0.7435
[9 empty] | 7 | 2 | 0.35647 | 0.35901 | 0.35878 | 0.3504 | 0.33841 | 18.51 | 1.2546 | 1.1327 | 0.98616 | 0.86813 | 0.79502
[9 empty] | 6 | 2 | 0.3088 | 0.31159 | 0.31682 | 0.30749 | 0.29924 | 16.17 | 1.3302 | 1.1987 | 1.0287 | 0.90117 | 0.81492
[9 empty] | 3 | 2 | 0.19412 | 0.20136 | 0.20667 | 0.20617 | 0.19955 | 15.03 | 1.6342 | 1.4278 | 1.1883 | 0.99916 | 0.88839
[9 empty] | 0 | 2 | 0.11589 | 0.12357 | 0.13054 | 0.13309 | 0.1305 | 14.39 | 2.1264 | 1.7515 | 1.3836 | 1.107 | 0.9567
[9 empty] | -3 | 2 | 0.065501 | 0.07319 | 0.08077 | 0.08444 | 0.083282 | 12.9 | 2.9415 | 2.239 | 1.6472 | 1.2431 | 1.0392
[9 empty] | -8 | 2 | 0.0209 | 0.02641 | 0.033655 | 0.03713 | 0.036452 | 11.74 | 5.6767 | 3.7119 | 2.2214 | 1.4872 | 1.1647
[9 empty] | -12 | 2 | 0.0069468 | 0.010748 | 0.015363 | 0.017353 | 0.017835 | 11.59 | 9.9795 | 5.389 | 2.7479 | 1.6693 | 1.2359
[9 empty] | -17 | 2 | 0.0018863 | 0.0035966 | 0.005688 | 0.0067855 | 0.0070966 | 12.79 | 16.7841 | 7.2338 | 3.2567 | 1.799 | 1.2808
[9 empty] | -20 | 2 | 0.0010729 | 0.0020777 | 0.003315 | 0.0039989 | 0.0046532 | 11.7 | 18.7071 | 7.7253 | 3.4401 | 1.8364 | 1.2876
[11 empty] | Vout @ -22 deg (V) | Vout @ 0 deg (V) | Vout @ 23 deg (V) | Vout @ 50 deg (V) | Vout @ 71 deg (V) |  | Rv @ 1K load, -22 C | Rv @ 1K load, 0 C | Rv @ 1K load, 23 C | Rv @ 1K load, 50 C | Rv @ 1K load,71 C
[9 empty] | 18 | 1 | 1.0751 | 1.0472 | 0.99885 | 0.9448 | 0.86469 | 22.35 | 0.64107 | 0.6182 | 0.59423 | 0.57258 | 0.56519
[9 empty] | 15 | 1 | 0.76947 | 0.75878 | 0.73328 | 0.69251 | 0.67022 | 20.41 | 0.76291 | 0.71587 | 0.66376 | 0.63237 | 0.60444
[9 empty] | 13 | 1 | 0.60408 | 0.60138 | 0.58519 | 0.56151 | 0.54185 | 19.57 | 0.84398 | 0.78592 | 0.72643 | 0.67515 | 0.64071
[9 empty] | 12 | 1 | 0.53381 | 0.53381 | 0.52023 | 0.50347 | 0.48668 | 19.48 | 0.88461 | 0.8232 | 0.75925 | 0.69954 | 0.66043
[9 empty] | 9 | 1 | 0.36252 | 0.36663 | 0.36491 | 0.35367 | 0.34435 | 17.15 | 1.0135 | 0.93361 | 0.84908 | 0.77015 | 0.72129
[9 empty] | 7 | 1 | 0.27414 | 0.27705 | 0.27896 | 0.27526 | 0.26793 | 18.09 | 1.116 | 1.0297 | 0.92031 | 0.82557 | 0.76513
[9 empty] | 6 | 1 | 0.23669 | 0.23962 | 0.24563 | 0.24064 | 0.2361 | 15.71 | 1.1724 | 1.0797 | 0.95322 | 0.85352 | 0.78384
[9 empty] | 3 | 1 | 0.14681 | 0.1528 | 0.15822 | 0.15954 | 0.15596 | 14.51 | 1.4028 | 1.2587 | 1.0823 | 0.93791 | 0.84788
[9 empty] | 0 | 1 | 0.085738 | 0.091987 | 0.098278 | 0.10166 | 0.10096 | 13.68 | 1.7887 | 1.5197 | 1.2471 | 1.0338 | 0.91085
[9 empty] | -3 | 1 | 0.046655 | 0.052783 | 0.059319 | 0.063357 | 0.063576 | 12.2 | 2.4689 | 1.9389 | 1.4831 | 1.1612 | 0.99062
[9 empty] | -8 | 1 | 0.01335 | 0.017455 | 0.023347 | 0.026878 | 0.027188 | 11.02 | 5.009 | 3.3212 | 2.0426 | 1.4087 | 1.1215
[9 empty] | -12 | 1 | 0.0039801 | 0.0065432 | 0.010123 | 0.012204 | 0.013088 | 11.07 | 9.4544 | 5.0689 | 2.6028 | 1.6086 | 1.2048
[9 empty] | -17 | 1 | 0.0010024 | 0.0020682 | 0.003577 | 0.0046627 | 0.0051447 | 12.5 | 16.6734 | 7.0288 | 3.1795 | 1.7643 | 1.2628
[9 empty] | -20 | 1 | 0.00055977 | 0.0011627 | 0.002058 | 0.0027223 | 0.0033553 | 11.4 | 18.845 | 7.6893 | 3.3814 | 1.8177 | 1.2797

## Thermal Comp by changing Rload
[14 empty] | Power Level | Load resistor | Temperature (Degrees C): 23 deg data is the Output Voltage measured. At beginning and end of test cycle. Data was copied from previous worksheet (Detector Prototype data average) |  |  |  |  | Diode Linearity (Lin=(delta Pin (dB)/log(V1/V2))
Pwr Lvl (dBm) | Load Resistor | Vout @ -22 deg (V) | Vout @ 0 deg (V) | Vout @ 23 deg (V) | Vout @ 50 deg (V) | Vout @ 71 deg (V) | Diode Linearity @ 5 K load | Rvideo @ 5K load, -22 deg C | Rvideo @ 5K load, 0 deg C | Rvideo @ 5K load, 23 deg C | Rvideo @ 5K load, 50 deg C | Rvideo @ 5K load,71 deg C |  | dBm | K ohms | Vout @ -22 deg (V) | Vout @ 0 deg (V) | Vout @ 23 deg (V) | Vout @ 50 deg (V) | Vout @ 71 deg (V) | 10=Square law 20= Linear Region >20= Saturation
18 | 5 | 1.5608 | 1.5054 | 1.4204 | 1.3306 | 1.2144 | 24.69 | 0.65177 | 0.62853 | 0.60546 | 0.58296 | 0.57226 |  | 18 | 10000 | 1.7643 | 1.6946 | 1.5924 | 1.4858 | 1.3534 | 26.28
15 | 5 | 1.1632 | 1.13 | 1.0737 | 0.99949 | 0.9556 | 22.46 | 0.83067 | 0.76099 | 0.68108 | 0.65509 | 0.6265 |  | 15 | 10000 | 1.3565 | 1.302 | 1.2243 | 1.1304 | 1.0753 | 23.96
13 | 5 | 0.93424 | 0.91569 | 0.87465 | 0.82393 | 0.78521 | 21.24 | 0.96157 | 0.86453 | 0.7754 | 0.70813 | 0.66098 |  | 13 | 10000 | 1.1139 | 1.074 | 1.0103 | 0.94062 | 0.88902 | 23.3
12 | 5 | 0.83394 | 0.82079 | 0.78479 | 0.74646 | 0.71061 | 21.15 | 1.0317 | 0.92866 | 0.83093 | 0.73156 | 0.68589 |  | 12 | 10000 | 1.006 | 0.97324 | 0.91521 | 0.85567 | 0.80809 | 22.66
9 | 5 | 0.5803 | 0.57936 | 0.56611 | 0.53662 | 0.5141 | 18.42 | 1.2892 | 1.1181 | 0.95948 | 0.83328 | 0.76467 |  | 9 | 10000 | 0.72992 | 0.70892 | 0.67475 | 0.62605 | 0.59272 | 19.95
7 | 5 | 0.44601 | 0.44644 | 0.44091 | 0.42483 | 0.40591 | 19.21 | 1.503 | 1.2979 | 1.0748 | 0.91417 | 0.82552 |  | 7 | 10000 | 0.58009 | 0.56232 | 0.53569 | 0.5025 | 0.47293 | 20.88
6 | 5 | 0.38829 | 0.38972 | 0.39111 | 0.37465 | 0.36021 | 16.87 | 1.6211 | 1.3934 | 1.1334 | 0.95267 | 0.84608 |  | 6 | 10000 | 0.51419 | 0.49833 | 0.47977 | 0.44603 | 0.42116 | 18.38
3 | 5 | 0.24905 | 0.25659 | 0.25969 | 0.25443 | 0.24315 | 15.74 | 2.0818 | 1.7249 | 1.3434 | 1.0757 | 0.92624 |  | 3 | 10000 | 0.35274 | 0.34511 | 0.32947 | 0.30917 | 0.28819 | 17.27
0 | 5 | 0.15313 | 0.1616 | 0.16746 | 0.1667 | 0.16059 | 15.34 | 2.8072 | 2.1714 | 1.5941 | 1.2014 | 1.0067 |  | 0 | 10000 | 0.2391 | 0.23178 | 0.22085 | 0.20676 | 0.19292 | 17.05
-3 | 5 | 0.090909 | 0.09968 | 0.10673 | 0.10785 | 0.10385 | 13.9 | 3.9011 | 2.7812 | 1.8999 | 1.3481 | 1.0932 |  | -3 | 10000 | 0.16184 | 0.15513 | 0.14729 | 0.13692 | 0.12656 | 15.5
-8 | 5 | 0.033452 | 0.040103 | 0.047433 | 0.049172 | 0.046468 | 12.66 | 6.9906 | 4.4041 | 2.4879 | 1.5829 | 1.2064 |  | -8 | 10000 | 0.080222 | 0.075426 | 0.071035 | 0.064739 | 0.05768 | 13.82
-12 | 5 | 0.012897 | 0.018069 | 0.022915 | 0.02363 | 0.023028 | 12.21 | 11.1314 | 5.9885 | 2.9579 | 1.7363 | 1.2656 |  | -12 | 10000 | 0.041609 | 0.03971 | 0.036471 | 0.031836 | 0.028856 | 12.91
-17 | 5 | 0.0039588 | 0.0066245 | 0.008927 | 0.0094284 | 0.009248 | 13.23 | 17.3749 | 7.5334 | 3.3735 | 1.8352 | 1.294 |  | -17 | 10000 | 0.017716 | 0.016605 | 0.01495 | 0.012889 | 0.011641 | 13.66
-20 | 5 | 0.0023192 | 0.0039077 | 0.005296 | 0.0055882 | 0.0060717 | 11.96 | 18.9493 | 7.9269 | 3.513 | 1.8631 | 1.2988 |  | -20 | 10000 | 0.011109 | 0.010103 | 0.009017 | 0.0076705 | 0.0076489 | 12.23
Calculated Vout as Load resistance is changed to temperature comp values | [6 empty] | Calculated drift (dB) @ new R load)
Pwr Lvl (dBm) | Diode Linearity @ 5 K load | Vout @ -22 deg (V) | Vout @ 0deg (V) | Vout @ 23 deg (V) | Vout @ 50 deg (V) | Vout @ 71 deg (V) | Drfit @ -22 deg (dB) | Drfit @ 0 deg (dB) | Drfit @23 deg (dB) | Drfit @ 50 deg (dB) | Drfit @ 71 deg (dB)
Detector Load resis (Kohms): |  | 15 | 7 | 5 | 5 | 8
18 | 24.69 | 1.6908 | 1.555 | 1.4204 | 1.3306 | 1.2631 | 1.8688 | 0.97069 | 0 | -0.7 | -1.2589
15 | 22.46 | 1.2853 | 1.1743 | 1.0775 | 0.99949 | 0.99724 | 1.7201 | 0.83894 | 0 | -0.73332 | -0.75528
13 | 21.24 | 1.0468 | 0.95596 | 0.87465 | 0.82393 | 0.82117 | 1.6574 | 0.81994 | 0 | -0.551 | -0.582
12 | 21.15 | 0.94128 | 0.85925 | 0.78479 | 0.74646 | 0.74428 | 1.6701 | 0.83257 | 0 | -0.46 | -0.4868
9 | 18.42 | 0.67215 | 0.61128 | 0.56611 | 0.53662 | 0.54101 | 1.3735 | 0.61402 | 0 | -0.428 | -0.36283
7 | 19.21 | 0.52725 | 0.47437 | 0.44091 | 0.42483 | 0.42869 | 1.4921 | 0.61028 | 0 | -0.31 | -0.23445
6 | 16.87 | 0.46403 | 0.4156 | 0.39111 | 0.37465 | 0.38088 | 1.2526 | 0.44504 | 0 | -0.315 | -0.19417
3 | 15.74 | 0.30975 | 0.27688 | 0.25969 | 0.25443 | 0.25828 | 1.2051 | 0.43826 | 0 | -0.14 | -0.037128
0 | 15.34 | 0.2014 | 0.17691 | 0.16746 | 0.1667 | 0.17136 | 1.2298 | 0.36581 | 0 | -0.03 | 0.15343
-3 | 13.9 | 0.12844 | 0.11102 | 0.10673 | 0.10785 | 0.11134 | 1.1173 | 0.23744 | 0 | 0.06255 | 0.25505
-8 | 12.66 | 0.05472 | 0.046298 | 0.047433 | 0.049172 | 0.050122 | 0.78577 | -0.13321 | 0 | 0.198 | 0.30313
-12 | 12.21 | 0.023885 | 0.021401 | 0.022915 | 0.02363 | 0.024915 | 0.21977 | -0.36247 | 0 | 0.163 | 0.4437
-17 | 13.23 | 0.0082081 | 0.007998 | 0.008927 | 0.0094284 | 0.010021 | -0.4824 | -0.6314 | 0 | 0.314 | 0.66397
-20 | 11.96 | 0.0049082 | 0.0047378 | 0.005296 | 0.0055882 | 0.0065806 | -0.39502 | -0.57851 | 0 | 0.279 | 1.128
Scratch pad: Change load resistance here and watch thermal drift change
[13 empty] | Load Resistance with Thermistors: |  |  | 15 | 7 | 5 | 5 | 8
[13 empty] | Operating Temperature (deg C): |  |  | -22 | 0 | 23 | 50 | 71